# MARSH FUNNEL VISCOSITY

## SCOPE:

This test is designed for routine quick measurements of fluid viscosity.

- PRINCIPLE:

The Marsh Funnel is conical shape, 6 inches in diameter at the top and 12 inches long with a capacity of 1500 cubic centimeters.  A 12-mesh screen covers half of the top and is designed to remove any foreign matter from the fluid.  The fluid runs through the fixed orifice at the end of the funnel which is 2 inches by 3/16 inches in size.

- PRECAUTIONS:

  - Check the calibration of the funnel with fresh water.  One quart (946mls) of fresh water at a temperature of 70 ± 5°F (21 ± 3°C) should out-flow the orifice in 26 ± 0.5 seconds.

  - Clean and dry the funnel and cup thoroughly after each use.

  - Do not blend or flatten the brass orifice in the bottom of the funnel.

- PROCEDURE:

  - Hold the funnel in an upright position with the index finger over the outlet.

  - Pour a freshly obtained sample of fluid to be tested through the screen until the fluid level reaches the bottom of the screen.

  - Remove the finger from the outlet and start the stopwatch.  Using the measuring cup, measure the time for the fluid to fill the one-quart (946 mls) mark of the cup.

  - Measure the temperature of the fluid in °F or °C.

  - Report the time to nearest second as the Marsh Funnel viscosity and record the temperature of the fluid.

- EQUIPMENT:

  - Marsh Funnel

  - Measuring Cup

  - Thermometer

  - Stopwatch